DESIGN OF THE ZEUS FORWARD TRACKING DETECTOR
Sherbak Y.L., Chernyshov N.N.
Kharkov National University of Radioelectronics
0930436635, chernyshov@kture.kharkov.ua
The design and construction of the ZEUS forward tracking detector is described. The detector consists of 3 large planar drift chambers mounted in the forward direction with respect to the proton beam and a smaller rear detector of the same type. Test beam measurements performed at DESY show a single wire spatial resolution of 105 μm and a dE/dx resolution of 8.8%. The resolution of the three dimensional track coordinates and slopes provided by one chamber are found to be 90μm and 2.7 mrad, respectively.

1 Introduction
The ZEUS detector was built for studying the collisions of 820 GeV protons on 30 GeV elec­trons in the HERA storage ring located at DESY. The main components of ZEUS are the inner tracking detectors, which are surrounded by a 1.8 T superconducting solenoid, a high resolution depleted uranium scintillator sandwich calorimeter and a muon detection system. Good track­ing in the forward region, i.e. in the proton beam direction, is of particular importance where jet topologies and the Lorentz boost result in high track densities. In the rear direction good tracking helps in identifying the scattered electron in low Bjorken x and low Q2 interactions. The inner tracking system was designed taking these considerations into account. It consists of: a vertex detector VXD (a high resolution cylindrical drift chamber), a central tracking detector CTD (a large cylindrical drift chamber), the forward tracking detector FTD (a set of three planar drift chambers) and the rear tracking detector RTD. The four planar drift chambers were added to the centra] tracking detector to improve the angular coverage and the track reconstruction at small and large polar angles where the wire orientation of the CTD becomes unfavorable. The forward tracking detector consists of three disk-like chambers FTD. each providing three dimensional track information. The FTD's are separated by t wo gaps of 210 mm each, which are used to accommodate 2 x 2 modules of the transition radiation detector TRD. In this paper, we report on the design and the construction of the forward tracking detector. Results on the chamber performance are presented, which were obtained using test beam data taken at DESY. In addition to the single wire spatial resolution and the resulting resolution of the three dimensional track parameters, the energy loss measurement (dE/dx) was studied.

2 Design
2.1 Chamber Design
The structure of the forward tracking detector as well as the internal structure of the individual chambers were designed to meet the following general requirements:
- The redundancy in the number of digitizings per drift cell has to be sufficient to resolve the left-right ambiguity already at the cell level and to cope with inefficiencies as well as spurious hits.
- Each chamber has to provide three independent stereo views of a track element to ease the reconstruction of three dimensional track elements by redundant information.
- The number of measured three-dimensional track elements (i.e. the number of chambers) has to exceed the minimum of two to keep the efficiency for track finding high and to ease linking tracks seen by the forward detector to tracks seen in the CTD and the vertex detector.

Table 1 - Parameters of the FTD's and RTD

	FTD1	FTD2	FTD3	RTD
distance from collision point (mm)	1220- 1374	1586-1740	1952-2106	1218-1372
overall
inner radius (mm)	124	124	144	161
outer radius (mm)	820	1030	1210	620
sensitive area
inner radius (mm)	180	180	200	220
outer radius (mm)	695	905	10S5	-195
angular acceptance (mrad)	140-470	122-489	105-489	175- 350
detection	8° -27°	7° -28°	6° -28°	10° -20°
number of cells	204	252	300	168
number of signal wires	1224	1512	1800	1008
number of potential wires	1428	1764	2100	1176
total number of channels	4536			1008
gas volume (m3)	0.25	0.40	0.60	0.1-1
total gas volume (m3)	1.39
radiation thickness
sensitive area	0.09 X0	0.09 Xo	0.09 X0	0.09 X0
outer rim	0.13 X0	0.13 Xo	0.13 Xo	0.13 X0

Each chamber consists of three independent layers of drift cells with fixed wire orienta­tions, rotated by 120° going from one layer to the next (fig. 1). Each drift cell has six sense wires. Thus one chamber measures three projections (u, v, w) of a track element with up to six digitizings per projection. The basic mechanical concept of the planar drift chambers was developed with the aim to minimize radiation thickness and, at the same time, to have a structure rigid enough to take all internal and external forces without distorting intolerably. The resulting design is summarized in table 1.

Figure 1- Basic chamber design FTD3 as an example

The chamber body is a self-supporting structure made of printed circuit boards. The load of the wire tension is taken by this body structure rather than by massive frames. Each cell has its own wire mounts, which are attached to the chamber body. External forces are introduced into a thin and hollow aluminium structure at the perimeter of the body, rigidly connected to it. Between the inner edge of the aluminium pieces and the beginning of the cells is a free space of about Δr = 80 mm. This space is needed to accommodate wire mounts, signal cables, and HV distributors. There is also a reinforcement structure at the inner (beampipe) hole of the chamber. It is made of thin pieces of glass fiber epoxy and supports cylindrical aluminium sheets, 0.5 mm thick, to close the chamber.

2.2 Cell Design
A cross section of the FTD(RTD) drift cell. The inner dimensions are Δz = 48 mm and Δu (Δv, Δtw) = 24 mm. Each cell contains six sense wires (30 μm diameter), spaced by 7 mm. and seven potential wires (121 μm diameter), perpendicular to the chamber axis (i.e. the beam direction). The potential wires lie in the median plane between the cathode planes. The sense wires are staggered with respect to this median plane by ±150 μm to allow a resolution of the left-right ambiguity already at the cell level. The distance of 12 mm between the median plane and the cathodes was chosen to be small enough
- to keep the maximum drift time (for extreme B-field conditions) smaller than 350 nsec (i.e. 3-4 bunch crossings), thus allowing the use of the FTD's for triggering,
- to cope with the problem of track finding in dense forward jets', and
- to reduce the effect of the strong and rather inhomogeneous magnetic field present in the FTD region on the calibration of the chambers.
Five field shaping strips on either side of the drift cell (2 mm wide and spaced by 2 mm) allow to produce a homogeneous electric drift field even at the edges. All planar electrodes, i.e. cathodes, field strips, and outer shielding, were made by printing and etching the Cu cladded boards. The potentials of the wires, cathodes, and field shaping strips were optimized in view of the following criteria:
- Minimal inhomogeneities of the drift field at the edges of the cell.
- Equal surface fields (gas gain) on all six sense wires.
- Surface field on potential wires below 20 kV/cm (to avoid whisker growth).
- Minimal electron drift to potential wires or field strips.
- Electrostatic deflections small compared to the gravitational sagging.
The initial optimization was done with help of the programs WIRCHA [1] and DRIFT [2] and fined tuned experimentally. The resulting set of potentials for a gas mixture of argon/ethane (50/50) at atmospheric pressure is shown in table 2. The corresponding drift field is about 1.25 kV/cm and the gas gain is about 5 x 104. The cathodes and field strips are electrically subdivided, such that the gas gain can be reduced in the region around the beam pipe in case of intolerably high backgrounds from beam gas interactions or synchrotron radiation.
Figure 2 shows the truncated mean distribution for 2 GeV electrons under an incidence angle of 15° obtained with total charge (a) and filtered charge (b) pulses. The curves are filled Gaussian distributions. We define the dE/dx resolution of an FTD chamber (Δ(dE/dx)) as the standard deviation of the distribution fitted to the measured truncated mean values.
The truncated mean follows an approximate Gaussian distribution with a a that depends both on the Landau distribution for the energy deposition (which at the same time depends on the characteristics of the chamber) and the number of samples used in its calculation. This a will be taken as the intrinsic component of A(dE/dx) and will be called σintr. To estimate it Gaussians were fitted to the truncated mean distributions obtained for the different conditions under study (different beam energy, different incidence angles) obtaining very similar σ s (within 10%). σintr was taken as the mean of all of them. Notice the wider distribution obtained if filtered pulses are used. A tuning of the digital filter parameters may improve this result. The second contribution to Δ(dEJdx) comes from our limited knowledge of the processes that occur when the particle crosses the chamber and in particular from uncertainties in the correction factors applied to the data. As they will translate into systematic deviations of the measured dE/dx from the real values, we call this contribution σsyst. The value taken for σsyst for the FTD2 is the rms of the distribution formed by:
- The observed deviations of the mean values of the truncated mean distributions, with respect to expectations, when varying the beam energy (Bethe-Bloch prediction) and incidence angle (constant behavior).
- The deviations from the constant value when varying the incidcnce position of the beam (vertical and horizontal chamber displacement).

Figure 2 - Truncated mean distribution for 2 GeV electrons under an incidence

The results were σsysttot = σsystfil = 6.0%, with a statistical error of ±1.2%. Unlike the intrinsic resolution σintr,the contribution σsyst obtained for the FTD2 cannot be extrapolated in a simple way to an expected σsyst for the three-chamber setup. We therefore estimate an upper limit for Δ(dE/dx) by adding in quadrature σsyst (FTD2) to σintr (3 - chambers) which gives

Δ(dE/dx)tot < 8.8% Δ(dE/dx)fil < 9.7%

3 Conclusions
In one chamber, the energy loss of a particle is sampled 18 times along its track. To reduce the effect of the long Landau tail in the determination of the mean energy loss per particle, a "truncated mean" was calculated, where the samples with the largest energy loss were not used. Only events with at least 15 recognized pulses with drift times larger than 50 ns and with a reconstructed particle track were considered for the following analysis. The amount of truncation for an optimum particle separation was estimated to be about 30% . We eliminated in this analysis only 17-20% of the samples (3 out of 15-18) because of the low number of reconstructed pulses per event that we had under certain conditions.

References
1 J. Fehlmann et al., WIRCHA program, ETH Zurich, 1984
2 J. Va’vra, L. Roberts, DRIFT program, SLAC, 1982